EXECUTIVE SUMMARY

(SEPTEMBER 29, 2008)

St. Boniface Hospital Research Centre

Winnipeg, MB, Canada

Michelle J. Alfa, PhD, FCCM

Nancy Olson, BSc

Louise Buelow-Smith, RN

Test 1:  Ward-based Washer Disinfectors & CPD based Washer Disinfectors

[Published data; Alfa MJ et al Simulated-use testing of bedpan and urinal washer disinfectors: Evaluation of C.difficile spore survival and cleaning efficacy.  Am J Infect Control 2008;36:5-11.]

Due to the increasing Clostridium difficile infections, the recent virulence of the microorganism, and knowing that bedpans are processed in washer disinfectors based on the ward or in the central processing department of hospitals, this study was to determine the cleaning efficacy and the C. difficile spore killing ability of such washer disinfectors.

During this study, the ArjoHuntleigh Tornado was evaluated as the ward-based washer disinfector.  The Steris Reliance 444 washer disinfector was the CPD-WD.

This testing utilized Ultra-Violet Visible Marker (UVM), Test Object Surgical Instrument (TOSI), Artificial Test Soil (ATS), and autoclaved stool/urine that was inoculated with 
C.difficile spores.

The results indicated that the CPD-WD was able not only to adequately clean the bedpans of organic soil, but also to completely inactivate the 6 logs of C. difficile spores placed in sealed ampules inside the CPD-WD.  The ward-based WD did not demonstrate adequate cleaning, and upon further investigation, it was determined that installation errors were responsible for the inadequate cleaning.

It was also determined that the thermal conditions of the Ward-based WD did not adequately inactivate C difficile spores.  This is not unexpected as a washer disinfector is designed to disinfect (thermal conditions of 85oC for 1 minute) and would not be adequate to kill a spore.

Conclusion:

The results of the study indicate that user testing of the efficacy of washer disinfectors is critical to ensure appropriate functionality.

Ward-based WD’s still provide an effective means of disposing of fecal waste from bedpans without generating aerosols.  Although the currently accepted thermal decontamination parameter for all bedpan WDs (ie, 80C for 1 minute) is appropriate for killing vegetative bacteria, these temperatures are not adequate to kill C difficile spores on bedpans.

The article indicated that manufacturers may be able to make adjustments to bedpan washers to achieve conditions that will effectively eliminate C. difficile spores.

Test 2 – Phase 1: [Data not yet published]

Baseline Testing – ArjoHuntleigh Tornado with Intensive Cycle

Based on the results and conclusions of Test 1 above, additional testing was undertaken to determine if the ArjoHuntleigh Tornado bedpan washer disinfector, using its current intensive program, is sufficient to eliminate C. difficile spores from bedpans or whether alterations to the cycle parameters are needed.

This testing utilized the same test methods outlined in Alfa et al (2008) consisting of autoclaved stool that was inoculated with C.difficile spores and dried onto bedpans to assess the cleaning cycle.  Cryovials containing the feces/C. difficile spore mix were used to test the C. difficile spore killing ability of the thermal conditions of the Tornado.

Conclusion:

Although the bedpans were visibly clean and achieved ~ a 4 Log10 reduction of spores on the surface of the bedpans, the Tornado intensive program was not able to completely eradicate spores from the bedpans.

The low reduction of spores in the cryovials (< 2 Log10 reduction in spore level) confirmed that the thermal conditions of the Tornado were not adequate to kill C.difficile spores efficiently.  This is not unexpected as the thermal conditions are 85oC for 1 minute in this cycle.

Test 2 – Phase 2: [Data not yet published]

ArjoHuntleigh Tornado with modified Intensive Cycle and specially formulated ArjoHuntleigh detergents

Based on the results and conclusions of Test 2 – Phase 1 above, additional testing was undertaken to determine if the ArjoHuntleigh Tornado bedpan washer disinfector, using a modified cleaning program and detergent, is sufficient to eliminate C. difficile spores from bedpans.

The same test method used in Test 2 Phase 1 was used for Phase 2.  The thermal temperatures in the phase were modified to 95oC for 300 seconds.  An enzymatic formulation and a new proprietary alkaline detergent formulation (SUREWash C.d.) were tested in conjunction with a modified Intensive cycle that included a higher temperature and longer wash time.

Conclusion:

Enzymatic detergent testing:

Although the thermal conditions of the ArjoHuntleigh Tornado are not sufficient to completely kill the C. difficile spores in the cryovials, the modified intensive cycle, used in conjunction with the Enzymatic detergent was successful in achieving complete surface disinfection.  However, the bedpans did not appear visibly clean despite the fact that no C. difficile spores were detected on the surface of the bedpans.  The use of the enzymatic detergent and modified intensive cycle would not be acceptable.

Alkaline detergent (SUREWash C.d.)

Although the thermal conditions of the ArjoHuntleigh Tornado are not sufficient to completely kill the C. difficile spores in the cryovials, the modified intensive cycle, used in conjunction with the proprietary alkaline detergent was successful in achieving bedpans that were visibly clean and had no detectable C.difficile spores; however, the volume of water used for the modified Intensive cycle was considered too excessive for routine use (~ 60 L per cycle).

Test 2 – Phase 3: [Data not yet published]

ArjoHuntleigh Tornado with regular Intensive Cycle and 45 mLs of  ArjoHuntleigh alkaline detergent (SUREWash C.d.) per cycle

Based on the results and conclusions of Test 2 – Phase 2 above, additional testing was undertaken to determine if the ArjoHuntleigh Tornado bedpan washer disinfector, using a regular Intensive cycle at 85oC for 60 seconds, and the new proprietary alkaline detergent (SUREWash C.D.) is sufficient to eliminate C. difficile spores from bedpans.

The same test method used in Test 2 Phase 1 was used for Phase 3.

Conclusion:

Alkaline detergent (SUREWash C.d.)

The thermal conditions of the ArjoHuntleigh Tornado are not sufficient to completely kill the C. difficile spores in the cryovials (assesses only thermal effect).  The regular Intensive cycle with 45 mLs of the proprietary alkaline detergent (SUREWash C.d.)  was successful in achieving bedpans that were visibly clean and had no detectable C.difficile spores.

Thus the combination of the proprietary alkaline detergent (SUREWash C.d.) and the regular intensive cycle thermal conditions provided a validated solution to the problem of adequate reprocessing of bedpans that takes into account the unique conditions needed to address C.difficile spores.